## data_dictionary
name | detected_dtype | detected_semantic_type | description | business_meaning | metrology_meaning | expected_range | missing_pct | unique_pct | potential_ml_usage | potential_preprocessing
id_instrument | object | free text | Unique identifier of the measuring instrument. | Primary key used to group all historical measurements of one physical instrument. | Instrument identity — the entity whose drift/anomaly is being tracked over time. | Non-null string, unique per physical instrument. | 0 | 2.87 | Grouping key for per-instrument feature engineering (trend, rolling stats); NOT a model feature itself (high cardinality identifier). | Normalize formatting (case, whitespace); verify 1:1 mapping with 'lab'.
lab | object | free text | Laboratory / instrument location label as recorded in the source Excel file. | Identifies which laboratory or production line the instrument belongs to. | Contextual grouping for control-chart analysis; not itself a metrology quantity. | Non-null string. | 0 | 4.82 | Categorical feature (encoded) if lab-specific effects are suspected. | Normalize text (accents, case, trailing year suffixes like '-2025') before using as a categorical feature.
annee | int64 | numeric | Calibration/verification campaign year. | Identifies which yearly campaign a measurement belongs to. | Time axis for drift analysis across calibration cycles. | Integer, plausibly between 2000 and 2035. | 0 | 0.78 | Base for time-since / campaign-count features; not used directly as a raw feature. | Validate range; combine with 'point' to build a proper chronological order.
point | object | numeric-like text (mostly convertible) | Measurement point identifier within a campaign (e.g. a calibration point number or a named position). | Distinguishes multiple measurement points taken during the same calibration/verification. | A single point on the instrument's calibration curve, or a named test position. | Expected numeric point index, but may contain text labels (observed in source data). | 0 | 3 | Ordering key within a campaign; needs type harmonization before any numeric use. | Harmonize mixed numeric/text values ('1' vs '1.0' vs 'au fond') into a single consistent representation.
consigne | float64 | numeric | Nominal/setpoint value the instrument is expected to indicate. | The target value used to assess how far the instrument's indication is from what it should read. | Nominal indication value, compared against the calibration reference. | Numeric, same physical unit and order of magnitude as 'valeur_reference'. | 34.29 | 4.04 | Feature for computing indication error (consigne vs valeur_reference). | Impute or flag missing values (frequently absent when the reference itself was used directly as the nominal value).
valeur_reference | float64 | numeric | Reference (traceable, etalon) value recorded during calibration/verification. | The 'ground truth' value the instrument is compared against. | Traceable reference value used to compute the instrument's error and conformity. | Numeric, instrument-category dependent. | 0.13 | 41.59 | Core input for drift/error computation; central regression target driver. | Outlier screening per instrument category; must never be imputed silently (a fabricated reference value would invalidate conformity).
incertitude | float64 | numeric | Measurement uncertainty associated with the reference value. | Quantifies the confidence interval around the recorded reference value. | Combined measurement uncertainty (ISO/IEC 17025), used to compute min/max bounds and to weight augmentation noise. | Strictly positive numeric value. | 0 | 20.73 | Feature 'uncertainty_ratio' (uncertainty / EMT width); used as noise scale in scientifically-grounded data augmentation. | Flag zero/negative/missing values as physically impossible.
EMT_basse | float64 | numeric | Lower bound of the Maximum Permissible Error (EMT) tolerance interval. | Regulatory/specification lower conformity limit for this instrument category. | Lower EMT bound; instrument is non-conformant if min falls below this value. | Numeric; must be strictly lower than EMT_haute. | 4.17 | 1.96 | Reference for 'distance_to_EMT' feature and for the deterministic rule engine (chapter 4 of the specification). | Flag rows where EMT_basse >= EMT_haute as impossible.
EMT_haute | float64 | numeric | Upper bound of the Maximum Permissible Error (EMT) tolerance interval. | Regulatory/specification upper conformity limit for this instrument category. | Upper EMT bound; instrument is non-conformant if max exceeds this value. | Numeric; must be strictly higher than EMT_basse. | 4.17 | 1.83 | Reference for 'distance_to_EMT' feature and rule engine. | Flag rows where EMT_haute <= EMT_basse as impossible.
min | float64 | numeric | Minimum bound of the measurement (typically valeur_reference − incertitude). | Lower edge of the plausible measurement interval used for the conformity decision. | Lower measurement bound compared against EMT_basse for conformity. | Numeric; must be lower than or equal to 'max'. | 0 | 70.53 | Direct input to the conformity rule; candidate regression feature. | Flag rows where min > max as impossible / corrupted source data.
max | float64 | numeric | Maximum bound of the measurement (typically valeur_reference + incertitude). | Upper edge of the plausible measurement interval used for the conformity decision. | Upper measurement bound compared against EMT_haute for conformity. | Numeric; must be greater than or equal to 'min'. | 0 | 71.32 | Direct input to the conformity rule; candidate regression feature. | Flag rows where max < min as impossible / corrupted source data.
conforme | object | categorical | Conformity verdict for this measurement (conforme / non conforme). | The quality decision actually recorded for this measurement point. | Ground truth conformity label, derivable from min/max vs EMT and usable to cross-validate the rule engine. | Categorical: {'conforme', 'non conforme'} (missing values observed in source data). | 13.82 | 0.26 | Candidate label to validate the deterministic rule engine and as an auxiliary classification target. | Normalize casing/spelling; investigate missing values (may indicate an unrecorded verdict, not a true unknown).